**RainbowSTORM Test Data**

## Install RainbowSTORM

## Download the RainbowSTORM plugin (Rainbow\_STORM.jar)

* 1. Copy the file into the Plugins subfolder of your ImageJ installation (e.g. “C:\Program Files\ImageJ\plugins”)
  2. Verify the successful installation of RainbowSTORM
  3. Restart ImageJ
  4. Locate RainbowSTORM under the Plugins menu
  5. Download and install the latest version of ThunderSTORM
  6. For ImageJ installations the Bioformats plugin is also needed to load the provided test datasets.

1. Access Help Screens
   1. Launch the RainbowSTORM Help Screen (from the ImageJ menu Plugins🡪RainbowSTORM🡪RainbowSTORM Help)
   2. Use hyperlinks to view help for each RainbowSTORM screen
   3. Help pages for each screen can be accessed by pressing the blue question mark icon on each screen.
   4. Close the Help Screens
2. Calibrate the System
   1. Download the calibration images (Calibration.tif)
      * Calibration Image Properties;
        + Calibration Source: Calibration Lamp
        + Dispersive Element: Grating
   2. Open ImageJ
   3. Drag the calibration images into the ImageJ window or load the calibration images
   4. Launch the RainbowSTORM Calibration Screen (from the ImageJ menu Plugins🡪RainbowSTORM🡪sSMLM Calibration)






Figure 1: RainbowSTORM Calibration Screen

* 1. Double-click on the line plot displayed in the uncalibrated spectrum graph to update the value of the current peak value text field starting with the first peak representing the spatial peak position). Click and drag to zoom in on an area.
  2. Press the 'Set' button to update the activated text field. Use the 'Previous' and 'Next' buttons to change which text field is activated.
  3. Select the spectral peaks for example the default wavelengths correspond to the spectral peaks at 487.7 nm, 546.5 nm and 611.6 nm indicated in Figure 1 below:




**Figure 2:** Example of spatial and spectral peaks for the example calibration image

* 1. The default settings for the calibration source (calibration light source) and the dispersive element(grating) are appropriate for these calibration images
  2. Press the 'Calibrate' button. The R-squared and RMSE values on the pixel versus wavelength plot to assess the calibration results.




**Figure 3:** Example of poor versus good fitting of calibration results

* 1. Save the resulting pixel positions, corresponding wavelength information, and fitting method
  2. Close the RainbowSTORM Calibration screen


1. Analyze the sSMLM Images
   1. Download the sSMLM images (Test\_A\_stack\_0107.tif)
   2. Drag the sSMLM images into the ImageJ window or load the sSMLM images
   3. Launch the RainbowSTORM sSMLM Analysis Screen (Plugins🡪 RainbowSTORM🡪sSMLM Analysis)

**Figure 4**: RainbowSTORM Analysis Screen

* 1. Set the camera parameters
     + Camera pixel size: 160 nm
     + Photoelectrons per A/D count: 4.6
     + Base level [A/D counts]: 200
     + EM Gain: 100
  2. Use the Rectangular Selection tool from ImageJ's toolbar to select the spatial image in the sSMLM image or manually input the parameters of the rectangle into the text fields (initial x position, initial y position, rectangle width, and rectangle height) on the sSMLM Analysis screen
  3. Press the 'Crop sSMLM Image' button on the sSMLM Analysis screen to separate the sSMLM images into spatial and spectral images




**Figure 5**: Example of the separating the spatial images from the spectra images


* 1. Optional-Press the 'Save Parameters' button to save the camera setup and cropping parameters (use an appropriate name to prevent overwriting other RainbowSTORM files)
  2. Select the Spatial Region (Cropped Region 1) and open ThunderSTORM

**Figure 6**: Example of analyzing the spatial images using ThunderSTORM

* 1. Make sure the camera settings used in ThunderSTORM and RainbowSTORM are the same
  2. Analyze the spatial images using ThunderSTORM (the default settings are appropriate)
  3. Export the results (don't overwrite any of the previously saved files)
     + Optional- deselect unwanted fields (Required ThunderSTORM fields: frame, x, y, intensity, sigma and uncertainty)
  4. Select the drift correction tab and select the cross-correlation default options. Check the saved to file text box and the ellipses to specify the filename. Press the apply button to apply and save the drift correction.



**Figure 7**: Example of saving ThunderSTORM drift correction information

* 1. Close ThunderSTORM
  2. Select the open RainbowSTORM screen
  3. Load ThunderSTORM results into RainbowSTORM
  4. Load Calibration Information into RainbowSTORM
  5. With the Automatic Background Subtraction' checkbox, press the 'Subtract' button
  6. Default settings are optimized for Far Red dyes e.g. Alexa Fluor 647
  7. Optional - Check the 'Remove Overlapping Spectra' checkbox to excluded localizations with overlapping spectra
  8. Optional - Preview the results of the spectroscopic analysis for individual localizations
  9. Press the 'Run Analysis' to calculate the results of the spectroscopic analysis for all localizations

**Figure 8:** RainbowSTORM Visualization Screen

* 1. Select a field (e.g. spectral sigma) from the histograms of the localizations for each RainbowSTORM field can be generated by selecting the field and pressing the "Show Histogram" button.
  2. Use the selected histogram to identify and a minimum and maximum value for the Filter Range. Input the values (e.g. spectral sigma 17 nm to 28 nm) and press the ‘Apply Filter’ button.
  3. Press the ‘Reset sSMLM Data’ button to remove filter results
  4. Update the Visualization settings:
     + Reconstruction method: Averaged Gaussian
     + Magnification: 5
  5. Load the drift correction file and press the correct drift button
  6. Press the ‘Calculate FRC’ to show results from Fourier Ring Correlation (FRC) Analysis
  7. Close FRC Results
  8. Use the ImageJ's rectangular selection tool is used to select a region of interest (ROI) in the current sSMLM reconstruction then pressing the "Restrict to ROI" button to select localizations within the ROI.
  9. Press the "Classify by Centroid" button launches the RainbowSTORM Classification screen where localizations can be grouped into color-coded image windows based on their spectral centroids.
     + Set three spectral windows (W1: 680 nm-689 nm, W2: 692 nm-698 nm and W3: 703 nm-709 nm)
     + Optional - Check the ‘Show Channel’ checkbox to display a reconstruction of each individual channel
  10. Press the ‘Save Spectroscopic Data’ button to save the current spectroscopic data.
  11. Close the Visualization screen.

1. Import previously saved data
   1. Launch the Import Screen
      * Select Plugins🡪 RainbowSTORM🡪sSMLM Analysis from the ImageJ menu
   2. Load the Data: Previously saved data files can be loaded into RainbowSTORM by selecting the file ending in "\_spec.csv".
   3. Set Camera Setup and Cropping Parameters: The camera setup parameters and the cropping can be manually input, or the previously saved parameters can be loaded.
   4. Press the ‘Visualize Data’ button to render the Lambda-colored image and load the data into the Visualization Screen.
2. Launch Help Screens